## Slide 1: Diagnostic Pressure Gauges 55.G3
Philip Andrew
Work done by
F4E
&
IPP Garching

## Slide 2: Content
ITER neutral pressure measurement
Diagnostic Pressure Gauges (DPG)
  Location
  Description
  R&D
Development in Gauge emission source
  Filament
  Indirect heating
  Emitter clamp
Summary

## Slide 3: Neutral gas pressure in ITER
Pressure of neutral gas between plasma and vacuum vessel
  Main source of gas: neutralization of plasma where it contacts divertor
  Gas pressure and flow affects impurity and He ash exhaust
  Numerical simulation of neutral pressure:
Log10 pn (Pa)
Pn (Pa)
1 Pa
5 Pa
< 0.01 Pa

## Slide 4: Neutral Pressure Measurement
ITER measurement requirements
  Locations, pressure range, response time

## Slide 5: DPG Distribution in ITER
6 DPG heads (1 spare) in each of 4 Divertor Cassettes (# 8, 23, 26 44)
Poloidal cross section of ITER
4 DPG heads (3 spares) in each of 4 Lower Ports (#4, 6, 12, 18)
6 DPG heads (5 spares) in each of 2 Equatorial Port Plugs (#1&1))

## Slide 6: Diagnostic Pressure Gauge: ionization gauge
Base plate
Filament supports Material Molybdenum
Filament Ø 0,8 different materials, different coatings Brazing Au82 Ni18
Ceramic insulators Brazing CB 10 Ag 64,8 Cu 25,2 Ti 10
Control electrode
Acceleration grid with different transparencies
Ion collector
Housing
Baffle ring
Baffle bridge
Thermocouple
Prototype produced by IPT-Albrecht

## Slide 7
ASDEX Pressure Gauge
with filament emitter
problem with hot filament:
25A
1900 K

## Slide 8
Acknowledgement: Houyin WANG & EAST team in Hefei, China
ASDEX Pressure Gauge with std filament
Removed from EAST tokamak
Filament over powered & yielded (cause not fully understood)
Extreme example of problem with hot filament:
ASDEX Pressure Gauge
with standard filament

## Slide 9: Main Technical Issues for implementation on ITER
Filament life
  Higher i ₓ B force
  Longer total “filament on” time
Extended pressure range
  Saturation effect at high pressures
Sensing ion current at 200m distance
  Pre-amplifiers typically used would be in excessive radiation environment
  Risk of cross talk over various junctions along the cable length
II
Ion Current vs Pressure

## Slide 10
Pressure Gauge with LaB6 source
Molybdenum posts
Graphite heaters
LaB6 crystal
Pressure Gauge with W filament
Tungsten Filament
Variant of ASDEX gauge developed by W7X team

## Slide 11: Emitter change for ITER Gauge
ZrC
Electron emitter
Graphite heaters
Material choices: ZrC expected to be more radiation resistant
Geometric choices: ZrC could run hotter than LaB6, need compliant emitter support

## Slide 12: Emitter stack and frame geometry optimization
Frame geometry → provide compression and remain elastic during stack expansion
Emitter geometry → reduce heating current and power providing stable performance
Stress distributions

## Slide 13: ZrC emitter with frame
frame
Prototype produced by IPT-Albrecht
Emitter assembled at IPP Garching

## Slide 14: ZrC emitter with frame
Prototype produced by IPT-Albrecht
Emitter assembled at IPP Garching

## Slide 15: Summary
Pressure gauges design modification after Preliminary Design Review
  W filament replaced by polycrystalline ZrC
  No more issue with Lorentz force on filament
  Electrical resistance heating of electrical feedthrough reduced greatly ( 25A -> ~ 8 A )
  Operation more stable and reproducible
Final Design Review is scheduled for May 2022
Experimental validation of final design is ongoing