## Slide 1: JT60-SA Thomson scattering status 01-04-2020
Roberto Pasqualotto
for the EU-TS team

## Slide 2: Thomson Scattering in JT-60SA
- Divertor TS (not shown) in the wish list

## Slide 3: TS diagnostics layout in JT60-SA
CORE&EDGE
CORE&EDGE
EDGE
CORE&EDGE
EDGE
CORE
CORE&EDGE
CORE
CORE&EDGE
installation

## Slide 4: Performance requirements
The core TS system (P2) shall measure electron temperature Te and density ne profiles
46 positions from R=2.6m to R=3.725m,
scattering volume length 15 mm ; nominal value from fiber bundle magnification
radial spatial resolution (points distance) 25 mm
dynamic range 0.1-30 keV
Laser pulse repetition rate will be 50 Hz.
The edge TS system (P1) shall measure electron temperature Te and density ne profiles
50 positions from R=3.7m to R=4.17m,
scattering volume length 5mm; nominal value  < 10 mm required incl. optics blur
radial spatial resolution (points distance) 25 mm at R<3.9 m, 5.5 mm at R>3.99 and 10 mm in between
dynamic range 0.01-10 keV.
Laser pulse repetition rate will be 100 Hz.
For both systems:
The accuracy is expected to be better than 10% for Te and 5% for ne, at ne = 1 x 1019 m-3 (from simulation)  Best Effort

## Slide 5: Status design & procurement prep.
Annex B for PA Signed in Jan 2020 to allow Fiber procurement
Fibers Fiber path reduced F4E tender launched 17 October, contract awarded on 23 March to Polymicro-Molex
Polychromators Specifications for F4E tender in Feb 2020: P1 spectral range from 660 nm F4E tender (negotiated procedure) Feb-Jun 2020
Optics edge TS (P1) Specifications for EF procurement in Dec 2019 Two stage contract: competitive design + manufacturing On hold since February: issue on EF reimbursement rules
Mechanics edge TS (P1) CAD design and FEM analyses nearly completed, preliminary negotiations with one manufacturer started

## Slide 6: PLANNING
Impact of COVID-19 on manufacturing ?

## Slide 7: Edge TS – mechanical structure
Collection optics on a retractable arm
Decoupled from cryostat
On supporting structure standing on floor
Same structure to install port plug
Collection optics
Retractable arm
Supporting structure
Vacuum window & shutter

## Slide 8: Edge TS – mechanical structure
Port plug fixed to insertion tube held by same trolley

## Slide 9: Edge TS – mechanical structure
Supporting structure assembled by bolted components to fit in the tight space
Installation procedure in progress

## Slide 10: Edge TS – mech. structure weight
Weight
  Port Plug: 650 Kg
  Insertion Arm: 405 Kg
  Trolley*: 2000 Kg
  Support structure: 2250 Kg
  * so heavy to carry the port plug
CoGravity
P.Plug+ inser.tube
CoGravity
P.Plug+ inser.tube+ trolley

## Slide 11: Edge TS – mech. structure FEM analysis
Total deformation(tip)
max 5mm
Safety Factor min 2.6
Acceptable port plug bending and safety factor before plastic deformation,
but significant weight increase

## Slide 12: Edge TS – port plug vacuum leak test
He inlet
Pressure gauge
For inter-groove He leaks
Check port plug vacuum tightness without a large vacuum chamber
Port Plug He leak tested by “tent procedure”
Port Plug with Lid filled with He, 1.2 bar over-pressure
Also O-rings interspace is monitored

## Slide 13: P1 port plug & window flange
Front plate
water cooling
Window screened by stabilizing plate
Port plug extends into vacuum vessel
Max use of available space
T<80°C

## Slide 14: P1 port plug: vacuum window shutter
Shutter components:
  Actuator
  Shaft and supports
  Threaded connection
  Pins
  Shutter
  Shutter support
8 cm
10 cm

## Slide 15: P1 collection optics – not rad hard
Specification for CfT:
Av. Transmission in 660-1070nm >0.7
Spatial resolution (90% energy) < 250 mm
NA on fibers < 0.22
Magnification: 3.6
Chief ray as normal as possible to image surface
Throughput shall be maximised (F/# ~ 2.3)
0.13 mm optical resolution
fiber bundle ~3x1.5 mm
Preliminary design
4/5 lenses + 2 mirrors
Curved & tilted image surface
+/- 0.5mm focus tolerance
Discussion with suppliers
Two contracts for competitive design (April?)
Manufacturer selection (July?)

## Slide 16: P1 collection optics – fiber bundles
Compromise required between NA mismatch and defocusing
Fiber bundles

## Slide 17: P1 collection optics & fiber holder
Fiber bundles array
Retractable arm
5cm
3cm
Optics block & fibers on same frame
Connected to retractable arm with hexapod system
Fiber holder adjustments: - Manual focusing & along laser beam - remote actuator for translation and rotation across laser beam for alignment
>30mm clearance to port plug walls
Fiber holder

## Slide 18: P1 collection optics & fiber holder
Mock-up to test robustness & adjustment of hexapod connections

## Slide 19: Optical Fibers: shorter path in torus hall
P1 oblique: -27m
P2 horizontal: -19m
Extra length 10 m in the torus hall is included.
P1 oblique: 83 m
P2 horizontal: 91 m
laser room
P2
P1
Old
New
First pit
Second pit

## Slide 20: Optical Fibers: F4E procurement
FBP low OH content: 2ppm
specification
dB/m@600nm
kGy FBP
0.1 0.2
1 0.4
10 0.9
100 2.4
SPECIFICATION
Silica core / fluorine silica cladding / polyimide coating (max T=150°C)
200/220/250 micron
0.22 NA
1100 km
Delivery: Jan-Jun 2021
Neutron dose: 0.13 MGy (1016 n/cm2)
wide market survey of rad hard fibers
incl. cost of mid OH & H2 loaded
0.9 MGy
0.1 MGy

## Slide 21: Filter polychromators
Same interf. filters + APD polychromator design for both core & edge systems: filter bandpass only difference; well mature design (dark noise 80  15 pe in 20 y)
6 cubicles, 16 polychromators each
Tech specs based on performance, no constraint of optical components, detector model and electronic design.
Both systems: 6 channels (5+Rayleigh); filters defined by TS signal simulations.
Acquisition system (QST): 12 bit, 1GHz, 1024 points/pulse (1 ms) 1 GB local memory: data read out after 100s plasma discharge
Filters transmission including polychromator transmission & APD QE
1
2
3
4
5
4
5
3
2
1
Edge TS

## Slide 22: Polychromators F4E CfT
CfT 27 Feb 2020
INFO-DAY 3 March
F4E tender: Negotiated procedure: - Step 1: Request to Participate Selection - Step 2: Invitation to Tender Negotiation

## Slide 23: Polychromators F4E CfT

## Slide 24: Polychromators F4E CfT
8 alignment fibers will be detected by modules similar to the polychromator channels
Installation and calibration on site will be performed by the supplier
Digital polychromator control only an option

## Slide 25: Edge TS simulation
Solid angle, scattering angle & length of each scattering volume vs R
Plasma light background adds to TS signal and contributes to noise: calculated for specific plasma profiles
2D Te(R,Z), ne(R,Z) and Zeff(R,Z) maps from 2D maps of poloidal flux and 1D Te and ne profiles
Plasma bremsstrahlung spectral emissivity integrated along each line of sight & wavelength integrated.
x2 enhancement factor to include line radiation and wall reflection
The plasma light signal was included as a noise source in Te and ne errors, in addition to shot noise and detector noise mostly due to preamplifier noise (dark rms noise ~30 pe).
Profiles prepared from CRONOS by J.Garcia and used by S. Fututani for JOREK input (pedestal profile shifted outwards)

## Slide 26: Edge TS simulation
ne (1019 m-3) | 0.87 | 0.6 | 0.37 | 0.18
dTe/Te (%) | 12 | 15 | 23 | 45

## Slide 27: Project challenges
Ref. to Sam presentation yesterday: my view of this project
Project team built on purpose from different labs: more effort but also opportunity
Negotiations with QST (Annex B): solution is a single central reference (F4E)
Lack of Engineering resources: support from F4E engineering team
PL has little knowledge of the subtleties and rigidity of Eurofusion rules, e.g. for reimbursement
PL has little knowledge and no control of own lab – Eurofusion relationship
PL has no knowledge and no control of partner lab – Eurofusion relationship
Weakness of initial budget estimate: best guess in haste at project kick off; cost increase also due to scope and constraints evolution
Lead time risk: mainly on testing / acceptance phase, to be well planned.