## Slide 1: The infrared imaging video bolometer at Wendelstein 7-X Paper rehearsal
G. Partesotti*, G. A. Wurden, F. Reimold, B. J. Peterson, F. Federici, K. Mukai, K. A. Siever, J. von Miller, D. Zhang, A. Demby, and the W7-X Team.

## Slide 2: I. Introduction
Measurements of plasma radiation are essential for power exhaust and physics studies
Bolometers are a consolidated diagnostic to measure plasma radiated power
Plasma radiation measure temperature rise calculate absorbed power
Resistive bolometers: measure electrical resistance of resistor thermally connected to absorber
Imaging bolometers: image temperature distribution on foil using high-resolution IR camera
InfraRed (imaging) Video Bolometer: IRVB
resistive
imaging

## Slide 3: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 4: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 5: II. The IRVB design at W7-X
Large (7cm 9cm) 5m-thick gold foil
C-coated on both sides with CVD
5mm 5mm square pinhole
Foil housed inside steel endoscope
Four Al-plated mirrors used for imaging IR radiation
IR camera placed at endoscope back (TH)

## Slide 6: II. The IRVB design at W7-X
Measure plasma radiated power on the upper divertor unit, module 51 (AEK51 port)
Wide collection angle of 40° poloidal 68° toroidal. Coverage in the toroidal range .
View limited to upper portion of plasma cross-section and low-iota section of divertor target plates
Define 2026 pixel grid on foil surface (total of 520 channels of 3.3mm 3.3mm size)

## Slide 7: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 8: III. Measurement procedure
              IRVB channels (pixels) are not independent
              Heat diffusion on the foil surface and black body radiation need to be accounted for
              The incident radiation power is calculated by solving a 2D heat transfer equation [1]
              Thermal properties of the foil need to be calibrated
              Three calibration coefficients:
              [1] B.J. Peterson et al. RSI 2000

## Slide 9: III. Measurement procedure
              Main steps in the IRVB data analysis:
              The measured counts data is mapped from the IR camera image to the foil pixels
              The counts are converted to temperature using camera conversion coefficients
              The 2D heat transfer equation is applied to calculate the absorbed power
is the bolometer measurement (line-integrated power)
Units of power density on foil surface (Wm-2)
Counts on IR camera
2
on foil view
3
on foil view
1
Counts on foil view

## Slide 10: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 11: IV. Interpolation to foil reference frame
              Foil is only partially imaged due to mirror magnification being too large
              Only one portion of the Cu frame is visible
              Data values on the foil surface need to be assigned to the correct bolometer channel

## Slide 12: IV. Interpolation to foil reference frame
              Interpolation from high-resolution camera view to bolometer pixel grid (20 26 channels)
              Effect of optics estimated using a calibration plate with regularly spaced holes
              Bivariate spline interpolation used to go from camera view to foil view (and vice versa)
              Interpolation accuracy gets worse near edge of the view
              A mask is used to select acceptable pixels.

## Slide 13: IV. Interpolation to foil reference frame
              Foil is only partially imaged due to mirror magnification being too large
              Only one portion of the Cu frame is visible
              Data values on the foil surface need to be assigned to the correct bolometer channel

## Slide 14: IV. Interpolation to foil reference frame
              Conversion counts-to-temperature using camera calibration
              Some features can be identified in the temperature distribution
              The features seem to follow a different thermal response
              Probably due to bad interpolation

## Slide 15: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 16: V. Calibration of the foil thermal properties
              2D heat transfer equation:
              Three coefficients need to be calibrated:
              Emissivity (): power absorption and radiative cooling
              Heat diffusivity () [m2s-1]: temperature gradients
              Heat sensitivity () [KW-1]: magnitude of thermal response

## Slide 17: V. Calibration of the foil thermal properties
              Foil calibration is carried out in the lab with laser irradiation experiments [1]
              Foil thermal response is measured using a high-resolution IR camera
              The combination of is found that reproduces laser power and time profile
              Accuracy of calibration method is tested on synthetic data (3% error)
              [1] F. Federici et al. RSI 2023

## Slide 18: V. Calibration of the foil thermal properties
              Laser is applied with square wave temporal profile
              Thermal properties are extracted from foil response
              Applied to ROI surrounding the laser spot
              Calibration is spatially resolved on the foil surface
              = (0.754 0.001)
              = (4.63 0.49)10−5 m2s−1, = (672 25) KW−1
              Missing proper camera calibration and settings
              Calibration of the 5m gold foil is still preliminary
              Calibration is being repeated for multiple foils with optimized setup

## Slide 19: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 20: VI. Experimental results Inference of the incident radiation power
              Example OP2.1 plasma discharge in standard configuration (XP:20230315.37)
              Constant applied power and plasma density
              Steady-state signal (except for LBO injection @ 5.5s)
              Incident power [Wm-2] is calculated by solving 2D heat transfer equation
Incident power

## Slide 21: VI. Experimental results Inference of the incident radiation power
              Non-uniform foil response (buildup of features in IRVB image)
              Probably arising from bad imaging and non-uniformity of foil calibration coefficients
              Features are estimated after heating power cut-off and subtracted
              Rescaled during the discharge using exponential fit
Incident power
Exponential fit
Residual power pattern

## Slide 22: VI. Experimental results Analysis of plasma radiation
              Incident power (after correction) is converted to radiance [Wm-2sr-1]
              ( [m] bolometer channel size, [m2sr] channel etendue)
              Radiance is converted to chord brightness (or emissivity) [Wm-3]
              ( [m] intersection length with plasma volume)
Corrected incident power
Radiance
Chord brightness

## Slide 23: VI. Experimental results Improved design
              Adjusted imaging mirrors
              Image entire surface area of the foil
              Higher uniformity of the foil sampling
              More coverage of the foil frame
OP2.1 IR camera image
OP2.2 IR camera image

## Slide 24: VI. Experimental results Improved design
              Substituted 5.0m gold foil for 2.5m platinum foil
              Thinner foil higher sensitivity (higher signal-to-noise ratio)
              Platinum lower diffusivity (smaller bolometer pixels)
              2mm 2mm bolometer pixels Higher spatial resolution.
OP2.1 measurement
OP2.2 measurement

## Slide 25: VI. Experimental results Improved design
              Early OP2.2 results confirm and extend OP2.1 observations
              Larger field-of-view and improved image quality
              Noise (NEPD) reduced from 3.93 Wm-2 (5.0m Au) to 1.99 Wm-2 (2.5m Pt)
              Toroidal asymmetry of plasma emissivity (factor 2 across 10 cm distance)
OP2.2 measurement
OP2.2 measurement

## Slide 26: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 27: Overview
Introduction
The IRVB design at W7-X
Measurement procedure
Interpolation to foil reference frame
Calibration of the foil thermal properties
Experimental results
Conclusions

## Slide 28: VII. Conclusions
IRVB diagnostic operational
Calibrated thermal properties of multiple bolometer foils (10% uniformity)
Remove non-uniformities in foil thermal response using fitting procedure
Divertor radiation pattern characterized by field-aligned stripes
Toroidal radiation asymmetry
New OP2.2 design with improved signal-to-noise ratio, field-of-view and resolution

## Slide 29: Acknowledgments
This work has been carried out within the framework of the EUROfusion Consortium, funded by the European Union via the Euratom Research and Training Programme (Grant Agreement No 101052200 — EUROfusion). Views and opinions expressed are however those of the author(s) only and do not necessarily reflect those of the European Union or the European Commission. Neither the European Union nor the European Commission can be held responsible for them.
This work was supported by the U.S. Department of Energy at Los Alamos National Laboratory, operated by Triad National Security LLC, for the National Nuclear Security Administration (Contract No. 89233218CNA000001).
This research was supported by the NINS-DAAD International Personal Exchange Program, JSPS Grant \# 22KK0039, NIFS Grant \# SIU004 and for B. J. Peterson the Alfried Krupp Wissenschaftskolleg Greifswald underwritten by the Stiftung Alfried Krupp Kolleg and the Alfried Krupp von Bohlen und Halbach-Stiftung.
G. A. Wurden, F. Reimold, B. J. Peterson, F. Federici, K. Mukai, K. A. Siever, J. von Miller, D. Zhang, A. Demby, and the W7-X Team

## Slide 30: Additional slides

## Slide 31: IRVB components
torus hall
foil (assembled)
parabolic mirrors
camera box (closed)
camera box (open)
endoscope front
foil + frame
endoscope cap

## Slide 32: Full camera view

## Slide 33: Calibration procedure
              Emissivity is calculated by comparing temperature to black-body reference
              are calibrated in two steps:

## Slide 34: Synthetic data
              Known power pattern applied to foil
              Forward calculate temperature distribution
              Foil thermal properties are known
              Add random noise fluctuations
              Simulation is resolved in time and space
Applied
on 5.0m Au
on 2.5m Pt

## Slide 35: Backward solution
              Use 2D heat diffusion equation to calculate power
              Compare with original input power
              Use synthetic data to test accuracy of calibration procedure
              3% error for 5.0m Au. 1% error for 2.5m Pt.
              Minimal 10mm 10mm calibration area for 5.0m Au (5mm 5mm for 2.5m Pt ).

## Slide 36: IRVB schematic

## Slide 37: IRVB geometrical parameters