## Slide 1: Progress in analysis of the JET LIBS data
Rongxing Yi et al.
Forschungszentrum Jülich GmbH, Institute of Fusion Energy and Nuclear Waste Management – Plasma Physics, Partner of the Trilateral Euregio Cluster (TEC), 52425 Jülich, Germany
PWIE, Juelich

## Slide 2: Outline
Background
In-vessel LIBS applications at JET
First results of JET-LIBS
Results in the LID-QMS region(Module 14)
Outlooks and Conclusions

## Slide 3: Background
Problem: Hydrogen isotope retention + erosion and co-deposition of wall material
JET-LID-QMS
Specific area
Fuel (high sensitive and quantitative)
JET-LIBS
Global
Fuel and PFM (low sensitive and quantitative in progress)
D-T 3

## Slide 4: In-vessel LIBS applications at JET
Optics Table
2.Littrow spectrometer: F#~4,resolution~30 pm(slit 100um), wavelength range 4 nm- Hα detection
1
2
3
4
4. PMT :monochromator +3 different filters- Hα detection
1. Echelle : F#10,resolution~70pm, 200-750 nm- T calculation
3. Compact C-T spectrometer: resolution 0.5 nm, 300-1100 nm- Layer information
Fiber bundle
Laser box
4 spectrometers
JET-LIBS setup
1
2
3
4

## Slide 5: In-situ LIBS applications at JET
Optics Table
MASCOT
H
D
W
W
JET-LIBS setup with remote handling arm
JET-LIBS setup
Comprehensive detection capability in Vessel

## Slide 6: First results of JET-LIBS
HFGC_MOD2_RH2W
2IW G1B
2IW G3A
2BW G4B
LBSRP_2W_RH
2ON G7A
2ON G8B
6 mm
Ablation craters on JET tiles
Position quite accurate
But profile not available

## Slide 7: First results of JET-LIBS
Plansee W
Ablation rate varies from the initial 30 pulses to 1000 pulses
Laser pulse number— real depth info
Ablation rate and reference sample
d~130 μm
Before we go into JET chamber, some test in lab give us crater info

## Slide 8: First results of JET-LIBS
FWHM of Ha is narrow enough to separate H/D after 8 μs. After 18 μs it reaches to 0.1 nm, which is still higher than the gap between T/D.
However, we can use voigt fitting to resolve them.
Spectral data
Different delay time

## Slide 9: First results of JET-LIBS
Read spectral data, check if there is significant (6 times larger than noise)W line @ 657.39nm (pixel 840+-5)
Use the W line do calibration, pixel difference between H D T, other W lines and W 657.39 nm is solid
Give the pixel position to each peak and do voigt fitting(solid and shared gaussian part, Lorentz part only shared by HDT )
Do the fitting directly with H. Set the original value for H peak position(526-531) , get the fixed value,
Save all data, but in red color when signal intensity <6N(only when signal >6N, it is regarded good signal)
yes
no
Spectral fitting
T changed in 2 weeks experiment
Wavelength 6 pixels deviation
No solid initial value for H
W line interference make it difficult to find the accurate wavelength for H
Problems:

## Slide 10: First results of JET-LIBS
4 typical spectra
Spectra with W and without W can both be fitted well.
Only when Ti peak appears, some unknow part make the spectral wings unfitted
Spectral fitting

## Slide 11: First results of JET-LIBS
Fuel retention in module 2 divertor(poloidal from HFGC-Tile8)
Too weak D

## Slide 12: First results of JET-LIBS
Absolute value
Due to the different delay times and different matrix, this absolute value didn’t make too much sense here

## Slide 13: First results of JET-LIBS
Layer information—Avantes spectrometer
HFGC
co-deposition: Be
D1
D2
D3
Laser pulses
Laser pulses
Laser pulses
Laser pulses
Intensity
Intensity
Intensity
Intensity
Edge effect
Thickness—distance

## Slide 14: First results of JET-LIBS
2D layer information for PFM– Be, W, Mo, C also H isotopes

## Slide 15: First results of JET-LIBS
Semi-quantitative analysis of the retained fuel composition
D concentration: Maximum of 0.5 at%,
slowly decreases to 0.1 at% by the 100th laser pulse.
D content in the initial 100 laser pulses~1.5×1017 D/cm2 (tile 8, assuming 179 nm/pulse ablation rate and ignoring the initial 7 empty pulses, which correspond to a 1.3 μm surface layer).
LOD: 0.07 at% (noise from pixel around 655 nm)
IU 667: 2 at% D in deposited W by NRA
As reference sample, W line 657.39 nm as the internal standard line

## Slide 16: Results in the LID-QMS region (Module 14)
The locations where LIBS and LID-QMS shot
In Module 14, Divertor area:
Tile 0 and 1, 4 tiles were measured by both LIBS and LID-QMS
The results can be compared somehow

## Slide 17: Results in the LID-QMS region (Module 14)
D empty position
1
10

## Slide 18: Results in the LID-QMS region (Module 14)
Highest D ratio
1
11

## Slide 19: Results in the LID-QMS region (Module 14)
1
6

## Slide 20: Outlook and Conclusions
  Conclusions:
  H and D peaks are well separated
  T is too weak to be detected, which means a very low concentration( baking performance is good)
  Ideal depth profile was obtained for gas and wall materials
  Semi-quantitative results were obtained for tile 8

## Slide 21: Outlooks and Conclusions
  Outlooks
  Calibration of echelle spectrometer for quantitative analysis
  Some unknow background in the H/D/T position to be studied
  Test JET-sample with LIBS-TALIF to measure T/D.

## Slide 22
Thanks for your attention!