## Slide 1: Update and Plans SPX
Ivo Classen
DIFFER

## Slide 2: WP-PWIE SPX: Plasma characterization, laser based diagnostic development
In 2025 SP X consists of 2 parts:
SP X.1: Atomic and molecular processes in attached/detached plasmas
SP X.2: Optimization of laser-based surface analysis diagnostics

## Slide 3: SP X.1
Deliverable ID | Beneficiary/ PM | Deliverable Title
D001 | DIFFER 8 | CARS results in Magnum-PSI. Flow velocity measurements with CIS and multispectral imaging using MANTIS on Magnum-PSI/UPP. (DIFFER)
D002 | FZJ 4 | VIS/VUV molecular and atomic spectroscopy on H and D in PSI-2. Comparison of MANTIS to hyperspectral camera in UPP. (FZJ)
D003 | CU 2 | VUV OES measurements in MAGNUM-PSI/UPP with focus on H/H2 (CU)
D004 | MPG 2 | Interpretation spectroscopy results using spectra simulation and (CR) modelling with YACORA and development VUV-OES (MPG)
Device | Beneficiary | Days | Related Deliverable
MAGNUM-PSI | DIFFER | 11 | D001, D002, D003, D004
UPP | DIFFER | 11 | D001, D002, D003, D004
PSI-2 | FZJ | 9 | D002
JULE-PSI | FZJ | 0 | D002

## Slide 4: D001: Diagnostic developments for plasma characterization in Magnum-PSI
SP X.1
D001
TALIF
CARS
MANTIS
CIS
MANTIS

## Slide 5: D001: TALIF for H measurements
SP X.1
D001
Publications in progress:
Comparison of Xenon- and Krypton-based TALIF Calibration Schemes for Hydrogen Density Determination
TALIF Diagnostic design for divertor relevant linear plasma devices Upgraded Pilot-PSI and Magnum-PSI
Measurements of atomic hydrogen density, temperature, and velocity in Magnum-PSI and UPP
Spectra above and below plasma column with two H populations (left). Average rotational velocity of H atoms in various B-fields (right) .
TALIF chord
Plasma column
Example acquisition from Magnum-PSI showing raw image (left), background-subtracted image (right)
1D H density profiles in Upgraded Pilot-PSI (UPP)

## Slide 6: D001: CARS for H2 measurements
SP X.1
D001
Last experimental campaign (Feb-March 2025) unsuccessful
Imperfect laser components; hotspots caused optical damage
Large distance at Magnum-PSI combined with implemented BOXCARS scheme worsened detection sensitivity
Technician visited last month to fix laser (6-month laser downtime)
CARS setup improvements still needed. Work performed last year:
Built and tested ‘Filtered CARS’ setup  Similar sensitivity to BOXCARS  less sensitive to large distances  Will be applied to Magnum-PSI
Tests performed with ‘polarization CARS’ during laser downtime  Use specifically polarized input lasers and add a polarization filter to remove background  Observed issues with ellipticity of output lasers due to coated components - Will not be tested during new Magnum-PSI experiments

## Slide 7: D001: MANTIS @ Magnum-PSI
SP X.1
D001
MANTIS: multispectral camera imaging[1]
  6 visible light cameras with unique filters, providing absolutely calibrated emissivities (photons/m3 sec)
  Hydrogen Balmer, H2 Fulcherband, Ar-I/II, Ne-I, C-II/III, N-II
  Inference of electron temperature and density, neutral hydrogen density, ionization rates, recombination rates, hydrogenic plasma-molecular interaction rates[2-5].
Aims:
  Test modifications to CRM through comparison against extensive set of Magnum-PSI diagnostics.
  Port improved parameter inference results to exhaust experiments equipped with MANTIS:
  TCV, MAST-U, AUG, W7-X, Magnum-PSI, RFX
First data next week
recorded emission x6
emissivity reconstruction x6
parameter estimate
Bayesian inference
(grid search or NN)
validation
inversion
(SART or NN)
collisional-radiative model
Thomson, TALIF
update
[1] A. Perek et al. Review of Scientific Instruments, 90(12):123514, 2019.
[2] A. Perek, et al. Nuclear Materials and Energy, 26, 2021.
[3] T.A. Wijkamp et al. Nuclear Fusion, 63(5):056003, 2023.
[4] A. Perek et al. Nuclear Fusion, 62(9):096012, 2022.
[5] B.L. Linehan et al. Nuclear Fusion, 63(3):036021.

## Slide 8: D001: Coherence imaging spectroscopy @ Magnum-PSI
SP X.1
D001
Coherence Imaging Spectroscopy is a camera-based polarization interferometer, enabling the acquisition of the 2D plasma flows.
Examples of use cases @ Magnum-PSI
  Visualize the pre-sheath plasma flows.
  Identify the entrainment of impurities with the main plasma flows.
  Quantify the momentum losses from attached to detached conditions.

## Slide 9: D002: VIS/VUV molecular and atomic spectroscopy on H and D in PSI-2
SP X.1
D002
Absolutely calibrated measurements of Balmer lines at PSI-2 using high resolution spectrometer and imaging spectrometer:
2 components of the Hα line were found (unshifted cold component (0.2-0.6 eV) and rotating hot component (3-4 eV, ionic nature doppler-shifted due to rotation of the plasma)).
Line-integrated emission profile of the Balmer lines, absolutely calibrated by integrating sphere. Spatial resolution across plasma column.
Experimental setup
z=0 mm
z=24 mm

## Slide 10: D002: Benchmarking plasma chemistry at PSI-2
SP X.1
D002
A Schumack PHYSICAL REVIEW E 78, 046405 2008
MAGNUM-PSI
Hβ line
PSI-2
D line, Klevarova V. Masterthesis (2015)
Two component hydrogen lines are observed at linear plasmas (MAGNUM-PSI, PSI-2, PISCES-A)
The most probable sources of the broad (rotating) component:
  e + H2+ → H + H (n=2,3,4,..) (dissociative recombination)
  e + H2+ → H + H (n=2,3,4,..) (dissociative excitation)
  (charge-exchange reaction is too weak)
The comparison of the Balmer lines emission with Yakora and AMJUEL
codes was performed using the new experimental data from imaging and high resolution spectrometers /Kotov V, EPS Conference P4.210 (2025)/.

## Slide 11
SP X.1
D002
D002: VIS/VUV molecular and atomic spectroscopy on H and D in PSI-2
Comparison/verification with CR modelling (Yakora, AMJUEL):
Dominant role of excitation due to dissociative recombination of the H2+ molecule and dissociation of molecules. Electron excitation and recombination negligible. Contribution of H2+ ions to the excitation of the H α line is ~ 60%. Contribution of H2+ ions to the Hβ line is ~80%. The second strong channel is the H2 dissociation.
The bands of molecular spectra of H and D measured at the PSI-2 in the VIS and VUV range are going to be analyzed until the end of 2025.
height, mm

## Slide 12: D004: VUV spectroscopy at UPP
SP X.1
D004
Friedl et al., Meas. Sci. Technol. 34 (2023) 055501
Fröhler-Bachus et al., JQSRT 259 (2021) 107427
Two devices applied:
VUV mini-spectrometer: 117-420 nm, 0.2 nm FWHM
VUV diode system with interference/edge filters: access to spectral range <113 nm
VUV spectrometer
VUV diode system
Ø 2.1 cm
Ø 13 cm  clipped to ~ 10 cm

## Slide 13: D004: VUV spectroscopy at UPP
SP X.1
D004
Friedl et al., Meas. Sci. Technol. 34 (2023) 055501
Fröhler-Bachus et al., JQSRT 259 (2021) 107427
Two devices applied:
VUV mini-spectrometer: 117-420 nm, 0.2 nm FWHM
VUV diode system with interference/edge filters: access to spectral range <113 nm
Absolutely calibrated spectra
Lya dominates spectrum
Detachment clearly to be seen in spectra (green vs. black):  H2 molecular systems (Lyman band, a-b continuum) decrease  Balmer series (n > 6) strongly increases, Balmer continuum appears instead of a-b continuum
Werner band + Ly series accessible with diode system

## Slide 14: D004: VUV spectroscopy at UPP
SP X.1
D004
Friedl et al., Meas. Sci. Technol. 34 (2023) 055501
Fröhler-Bachus et al., JQSRT 259 (2021) 107427
Two devices applied:
VUV mini-spectrometer: 117-420 nm, 0.2 nm FWHM
VUV diode system with interference/edge filters: access to spectral range <113 nm
Absolutely calibrated spectra
Lya dominates spectrum
Detachment clearly to be seen in spectra (green vs. black):  H2 molecular systems (Lyman band, b-a continuum) decrease  Balmer series (n > 6) strongly increases, Balmer continuum appears instead of b-a continuum
Werner band + Ly series accessible with diode system
Detached regime dominated by atomic radiation  Lyman series leads to high flux of high-energy photons (> 12 eV) even under detached conditions!
VUV applicability demonstrated  Insights into molecular radiation and recombination t.b.d.

## Slide 15: SP X.2
Deliverable ID | Deliverable Title
D001 CEA | Comparison of ps vs. ns LIBS: absolute content and composition. Comparison SP vs. DP LIBS, or alternative LIBS signal enhancement methods: absolute fuel content in W samples and composition (CEA)
D002 DIFFER | In-situ LIBS and NRA in Magnum-PSI. Reference measurements of outgassing and retention after D plasma loading: absolute content and composition in W and reference samples (DIFFER)
D003 ENEA | Comparison SP vs. DP LIBS, or alternative LIBS signal enhancement methods: absolute fuel content in W samples and composition / (CF)-LIBS results on H-D loaded samples and surface modifications. Investigation of the possibility to detect He with LIBS on He-loaded samples at ambient pressure and RT (ENEA)
D004 FZJ | Comparison of ps vs. ns LIBS: absolute content and composition / Comparison SP vs. DP LIBS, or alternative LIBS signal enhancement methods: absolute fuel content in W samples and composition / Reference measurements of outgassing, recycling, and retention after D plasma loading: absolute content and composition in damaged and undamaged Wsamples / LAMIS measurements and identfication of C and N containing molecules /Explore TALIF and LIBS for H isotope identification (FZJ)
D005 IPPLM | Report on LIBS analysis by application of machine learning algorithm: focus on development of the models for chemical contents estimation and reconstruction of experimental signals. (IPPLM)
D006 CU | ps/ns LIBS analysis of B, W, (Be) containing coatings. CF quantification and depth profile analysis. Re-excitation of weak ps laser induced plasma by 2nd ns laser pulse (DP LIBS ). Analysis of JET and ITER related samples. He-loaded and deuterium-exposed coatings with differing surface morphologies. Ablation rate analysis. (CU)
D007 UT | Analysis of B containing coatings and/or W coatings with (CF)-LIBS: absolute content and composition / (CF)-LIBS results He loaded samples and surface modifications / In-situ LIBS and NRA in Magnum-PSI: Reference measurements of (LIBS induced and natural) outgassing and retention after D plasma loading: absolute content and composition in W based samples. (UT)
D008 ISSP UL | Comparison of ps vs. ns LIBS: absolute content and composition / Analysis of B containing coatings and/or W coatings with (CF)-LIBS: absolute content and composition (ISSP-UL)
D009 VTT | LIBS measurements on W and Be containing reference layers in support of the SP E.1 activities (VTT)

## Slide 16: SP X.2
Device | Beneficiary | Days | Related Deliverable
MAGNUM-PSI | DIFFER | 7 | D002, D007
PSI-2 | FZJ | 15 | D004
Accelerator | DIFFER | 12 | D002

## Slide 17
LIBS developments
of the CORIA-CEA group in 2025
Floriane Montupet-Leblond, Elodie Bernard, Stéphane Vartanian, Marc Missirlian, Christian Grisolia
Arnaud Bultel‡, Vincent Morel, Aurélien Favre, Benjamin Quevreux
CORIA, UMR CNRS 6614, Normandie Université, 76801 Saint-Etienne du Rouvray cedex, France
‡arnaud.bultel@coria.fr
SP X.2
D001

## Slide 18
Temperature (K)
t0 = 10 ns
Electron density (m-3)
t0 = 10 ns
On Eurofer97 containing
Fe, Cr, Cu, W, Na, Mn, V measured by XRF
Plasma thickness (mm)
t0 = 10 ns
Saha-Boltzmann plot
Validation of the MERLIN code (spectra calculations)
Spectral radiance
relative difference
Spectral radiance
A. Favre, A. Bultel, V. Morel et al. JQSRT 330 (2025) 109222
https://doi.org/10.1016/j.jqsrt.2024.109222
At time t = 950 ns:
experimental spectrum
reconstruction by MERLIN
blackbody radiation
SP X.2
D001

## Slide 19
FLATEN
DENSE + ReLU x5
Multi-layer perceptrons: decision
DENSE + ReLU + Softmax
DENSE + ReLU + DENSE
DENSE + ReLU + DENSE
105 spectra calculated by MERLIN with 12 parameters:
mole fraction in Ag, Ar, B, C, Cu, H, Mo, Ni, W
around the wavelength
at
related to LIBS measurements performed in WEST
Coupling between Machine Learning and MERLIN for real time CF-LIBS
A. Favre, V. Morel… A. Bultel. Spectrochim. Acta Part B 224 (2025) 107082
https://dx.doi.org/10.1016/j.sab.2024.107082
x 6
CONVOLUTION
RESIDUAL BLOC x4
MAX POOL
NORMALISATION
ReLU
CONVOLUTION
STACK
Convolutional Neural Network: information extraction
Input
MERLIN
(wavelength, spectral radiance)
A. Favre, A. Bultel… V. Morel et al. Phys. Scr. 99 (2024) 035609
https://doi.org/10.1088/1402-4896/ad2826
Learning  Loss function:
where is the prediction, the target and the prediction confidence on the parameter
Best prediction!
Tests for the quantification:
Outputs:
9 mole fractions
12 x
Lowest predicted value for H:
Real time test of the whole procedure (1000 times):
Writing-reading of a spectrum
Calibration
Image treatment
Network response
At 10 Hz: analysis at each laser shot
SP X.2
D001

## Slide 20
Nd:YAG laser source
Power supply
Tokamak window
View system
Movable lens
Periscope
Plasma light collection system
Preparation of a LIBS device within WEST: the LIBS4FUSION project

## Slide 21: D002: LIBS on Magnum
SP X.2
D002
D002 is closely linked to D007
New LIBS campaign on Magnum in December 2025
Comparison LIBS vs NRA
Study outgassing during ns LIBS by measuring NRA depth profiles in LIBS craters of increasing depth (vary number of laser pulses)
Data successfully taken, analysis early 2026
LIBS setup at TEAC (with Jordy [DIFFER] and Jasper [UT]
Ion Beam Facility for NRA
Exposed samples in TEAC

## Slide 22: D003: Optimization of the LIBS technique in Air, He, and Ar
SP X.2
D003

## Slide 23: D003: Optimization of the LIBS technique in Air, He, and Ar
SP X.2
D003
The use of He and/or Ar as background gases allows for the separation of the Hα and Dα emission lines of hydrogen and deuterium with higher resolution than traditional LIBS in air, as soon as the acquisition parameters (ICCD camera gate delay and width) have been optimized.
The simulated Tα spectral emission of Tritium, having an intensity comparable to the Dα emission of deuterium allows to conclude that the detection of both lines requires a system with a high resolution be clearly detected, but the detection is feasible with post-measurement fitting procedure.

## Slide 24: D004: Cross comparison LIA-QMS and NRA
SP X.2
D004
General Information
Pre-damaged tungsten material could have enhanced reservoir for hydrogen isotopes
Depth profile of the D content in near surface damaged W sample can be measured by LIA-QMS
Task and questions to be addressed
Cross comparison of D content depth profiles measured by LIA-QMS and NRA
Approach
W sample damaged with the same dpa doses are loaded in PSI-2 with deuterium
LIA-QMS(=355 nm, 35 ps, 5 mJ, spot  0.46 mm) is applied for each sample
Calibration with H2 and D2 calibration leaks and , in addition, sample PlaQIII-1 total D content used for scaling factor measurement
Results
QMS HD+ signal (m/z=3) about 10 times larger D2+ signal (m/Z=4)  strong interaction of LIA released D with wall  reference sample with known content D is required
Usage H2 and D2 calibration leaks in combination with reference sample provided accuracy of the D retention measurement about 20%
LIA-QMS depth profile almost coincides with NRA profile except first 350 nm due to , probaly, thermal effect. Measured sensitivity of LIA-QMS < 0.05 at%.
Fig. 1. Reference sample QMS signals
Fig. 2. Test sampe depth profile

## Slide 25: D004: Investigation of hydrogen adsorption dynamics by LIBS
SP X.2
D004
General Information
Adsorption of the background hydrogen may influence on LIBS measurement in vacuum
Hydrogen adsorbed on tungsten surface can be used as LIBS calibration sample
Task and questions to be addressed
Investigation of hydrogen adsorption dynamics by LIBS at different background pressures
Approach
Annealed tungsten sample without H content were used
Single pulse LIBS (=355nm, 35 ps, 26mJ) in LIBS in 10-7 – 10-6mbar vacuum
5 laser pulses at the same laser spot were used to clean surface before start of measurements
H line was used for hydrogen detection in LIBS plasma
Results
Plasma continuum and tugsten spectral lines measured at 0.5 s time interal between laser pulses ( H line is invisible ) was used for background substraction from the spectra measured with longer intervals
LIBS H spectral line has two component blue and red shifted due to Dopler effect.
Total H intensity is increased with time interval for short intervals until it is saturated
Time interval reqired for saturation decreases with pressure increase
Fig. 1. LIBS H spectrum
Fig. 2. LIBS intensity vs adsorption time

## Slide 26: D004: Explore TALIF and LIBS for H isotope identification
SP X.2
D004
General Information
Spectral interference by W, Ti … and hydrogen isotopes reduce SNR with increase of delay time
Combination of LIBS plasma with TALIF will improve detection limit for hydrogen isotopes measurement
Task and questions to be addressed
Feasibility study of hydrogen detection by LIBS-TALIF
Approach
Tungsten sample with natural H content were used for LIBS-TALIF in air
Single pulse LIBS (=355nm, 35 ps, 10mJ) in air atmospheric pressure
OPO (Optical Parametric Oscillator) laser with =205.14nm, <0.042nm, =3-5ns, E=2mJ
OPO laser beam was ortoganal to =355nm laser beam
Results
First measurements were performed in air
There is no W line interference at all
H and D lines can be excited separately
Low noise in comparison with LIBS spectra
spectrometer
Fig. 1. LIBS-TALIF set-up
Fig. 2. TALIF spectrum
H
Gate 100 ns
Delay 20 s

## Slide 27: D005: ML models development at the IPPLM – autoencoder models for experimental data
SP X.2
D005
Spectra provided by Rongxing Yi, FZJ Juelich
Nd:YAG laser (5 ns, 50 mJ, 1064 nm)
Pure elemental samples of Al, Cr, Fe, Mo, Ni, W, and Zn
200 - 900 nm range, 26 000 wavelengths sampled
Linear combination of spectra simulating varying composition

## Slide 28: D005: ML models development at the IPPLM – results
SP X.2
D005
Autoencoder as Compression Algorithm
Compression ration of 170
For reference:
zip: 2-10
JPEG: 10-20
MP4: 50
Hidden | Latent | Regularizer | Activation | MSE | CosSim | Relative Error
1621 | 195 | 0.0 | leaky ReLU | 0.021 | 0.23 | 6.4%
1377 | 136 | 0.0012 | tanh | 0.20 | 0.041 | 11%
Reconstruction
With autoencoder:
MSE = 2.3 x 10-5
RMSE = 0.31%
R2 = 0.9993
Training time = 30 s
Without autoencoder:
MSE = 1.7 x 10-5
RMSE = 0.27%
R2 = 0.9993
Training time = 20 min
With the autoencoders the accuracy is insignificantly lower, but still excellent and the computative cost is substantially lower.

## Slide 29: D006: Single pulse ps LIBS for laser ablation- plasma characterization
SP X.2
D006
  • Pure W was firstly studied by ps LIBS (250 ps, 2 mJ) , where suitable W lines were identified and selected (depending on LIP parameters such as fluence).
  • For the BP lines selection, we used only lines not affected by self-absorption and interference (e.g. green markers in the figure are the suitable lines).
  • The evaluated Te of ps LIBS plasma generated by W ablation in air at atmospheric pressure were: 0.44 eV (for 80 ns / 40 ns) and 0.41 eV (for 500 ns / 250 ns). The observe Te is lower in early plasma afterglow comparing to our previous ps LIBS plasma measurements of W and Mo using 30 ps laser (pulse energy 0.3 mJ), where the fluency was similar

## Slide 30: D006: Double Pulse ps-LIBS for depth analysis
SP X.2
D006
  • Investigated D retention in W based coated materials with different surface morphologies, exposed to D plasma in the Magnum-PSI device.
  • Employed a dual-laser LIBS setup: a ps laser (250 ps, 2 mJ) for gentle ablation and a secondary ns Nd:YAG laser (8 ns, 40 mJ, 100 ns delay) under 200 mbar Ar to enhance plasma excitation and D-line detection, with spectra recorded using an Andor ME5000 echelle spectrometer and iStar DH743 iCCD (5 ns resolution). This was done after the single pulse ps LIBS studies.
  • The depth profiles suggest a thin near-surface D/H deposition layer of about 100 nm, but the depth values carry large uncertainty. The spot size of ~40– 60 μm was consistent and correct across measurements
  • Clear depth profiles were obtained, showing morphology-dependent differences in retention behaviour between compact, porous, and nanocolumnar structures.

## Slide 31: D007: In-situ LIBS and NRA in Magnum-PSI: LIBS induced outgassing
SP X.2
D007
DIFFER, UT, CU
Plans (Dec. 2025): consecutive laser shots and NRA measurements at the same position
Laser pulse
Material
Ablation
Outgassing
Both LIBS induced outgassing and diffusion observed
LIBS ablation 150-200 nm
Large LIBS crater to fit NRA!
NRA area
LIBS crater
633
317
950
1266
1583
1899
2216
0
101
100
10-1
10-2
5000
0
10000
15000
20000
25000
Depth (nm)
Depth (1015 D/cm2)
D concentration (at.%)
p-WO coating
NRA
Before LIBS shot
After LIBS shot

## Slide 32: D007: In-situ LIBS and NRA in Magnum-PSI: D retention after loading
SP X.2
D007
Chart: LIBS shot nr.
Chart: Depth (nm)
| 0 | 398.88 | 398.88 | 842.07999999999993 | 842.07999999999993 | 1152.32 | 1152.32 | 2000
 | 4 | 4 | 3 | 3 | 1.5 | 1.5 | 0 | 0
Chart: c-W thickness
| 1 μm | 1.5 μm | 2 μm
NRA | 183 | 204 | 380
LIBS | 125.4545 | 222.7273 | 410
Compact 1 µm, 1.5 µm and 2 µm tungsten (c-W) coatings on Mo substrate
LIBS spectra show a poorly defined interface between W coating and Mo substrate
SEM/EDX and RBS suggest mixing of W and Mo after LIBS shots
NRA/RBS shows 4-5 at. % D on c-W surface and slow D reduction to negligible value inside the c-W coating
D retention increases for thicker c-W coatings and the sum of LIBS D intensity follows the trend
W/Mo interface?
DIFFER, UT, CU

## Slide 33: D008: ps-LIBS on Boron
SP X.2
D008
ps-laser LIBS has been applied to the magnetron-sputtered B coatings on W substrate
Samples: 30…50 nm B coating was applied on W substrate by magnetron sputtering at the ISSP UL (Laboratory of Thin films)
Elemental depth profiling
White light profilometry procedure has been obtimised to obtain 3-D profiles of the laser ablation craters
Elemental depth profiling shows that it was possible to detect B singal only at the first pulse. The is due to the fact that B detection has required increased laser pulse energy, and the best ablation rate achieved now is 80 nm/pulse @30 mJ, according to the profilometry data
Improving of ablation rate for B detection is planned early next year

## Slide 34: Call for participation 2026 / 2027 Preliminary outcome
2026 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 | TASK | 4 | CEA | CU | DIFFER | ENEA | FZJ | IPP_LM | ISPP_UL | MPG | NCSRD | RBI | UT | VR | VTT | 
SPX | X1 | Atomic and molecular data expert |  |  |  |  | 2 | 0 |  | 2 |  |  |  |  |  | 
 | X2 | Laser-induced ablation and spectroscopy expert | 3 | 5 | 2 |  | 8 | 5 | 3 | 3 |  |  | 5 |  | 2 | 
 | X2 | Surface-analysis expert | 0 |  | 2 |  | 2 |  |  |  | 0 | 0 |  | 0 | 0 | 
 | X1 | Spectroscopy expert | 0 |  | 7 |  | 4 |  |  | 1 |  |  |  |  |  | 
2027 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
SPX | X1 | Atomic and molecular data expert |  |  |  |  | 2 | 0 |  | 2 |  |  |  |  |  | 
 | X2 | Laser-induced ablation and spectroscopy expert | 3 | 6 | 2 | 3 | 8 | 5 | 3 | 3 |  |  | 6 |  | 2 | 
 | X2 | Surface-analysis expert | 0 |  | 2 |  | 2 |  |  |  | 0 | 0 |  | 0 | 0 | 
 | X1 | Spectroscopy expert | 0 |  | 7 |  | 4 |  |  | 1 |  |  |  |  |  | 

## Slide 35: Thank you for your attention 3 Highlight talks: