## Slide 1: Introduction to LID-QMS diagnostic on JET, including type of data available, progress on results
Miroslaw Zlobinski et al.
Forschungszentrum Jülich GmbH, Institute of Fusion Energy and Nuclear Waste Management – Plasma Physics, Partner of the Trilateral Euregio Cluster (TEC), 52425 Jülich, Germany
Joint LID-QMS/LIBS meeting, 9.12.2025, Frascati

## Slide 2: Acknowledgements
Miroslaw Zlobinski1*, A. Huber1, G. Sergienko1, I. Jepu2, C. Rowley2, A. Widdowson2, R. Ellis2, D. Kos2, I. Coffey3, M. Fortune2, D. Kinna2, M. Beldishevski2, L. Laguardia4, G. Gervasini4, A. Krimmer1, H. T. Lambertz1, A. Terra1, S. Brezinsek1,5, T. Dittmar1, M. Flebbe1, R. Yi1, R. Rayaprolu1, J. Figueiredo6,7, P. Blatchford2, S. Silburn2, E. Tsitrone8, E. Joffrin8, K. Krieger9, Y. Corre8, A. Hakola10, J. Likonen10, THE EUROFUSION TOKAMAK EXPLOITATION TEAM17, THE JET CONTRIBUTORS18
1 Forschungszentrum Jülich GmbH, Jülich, Germany 2 United Kingdom Atomic Energy Authority, Culham Campus, Abingdon, United Kingdom 3 Queen’s University Belfast, Belfast, Northern Ireland, UK 5 Heinrich-Heine-Universität Düsseldorf, Düsseldorf, Germany 6 EUROfusion Programme Management Unit, Garching, Germany 7 Instituto de Plasmas e Fusão Nuclear, Instituto Superior Técnico, Universidade de Lisboa, Portugal 8 CEA, Institute for Research on Fusion by Magnetic confinement, St-Paul-lez-Durance, France 9 Max–Planck-Institut für Plasmaphysik, Garching, Germany 10 VTT Technical Research Centre of Finland Ltd, Espoo, Finland 17 See the author list of E. Joffrin et al., Nuclear Fusion 64 (2024) 112019 , doi:10.1088/1741-4326/ad2be4 18 See the author list of C.F. Maggi et al., Nuclear Fusion 64 (2024) 112012, doi:10.1088/1741-4326/ad3e16 *Email: m.zlobinski@fz-juelich.de

## Slide 3: Outline
Motivation
The LID-QMS method
Diagnostic Setup at JET
QMS Calibration
Different Evaluation Methods
Open Issues
in situ Applications at JET
Comparison possibilities with LIBS

## Slide 4: Motivation
Ref: S. Brezinsek et al., Nucl. Fusion 55 (2015) 063021, doi:10.1088/0029-5515/55/6/063021
Problem: Hydrogen isotope retention at the wall
Risk of high tritium release in case of large leak (e.g. ITER limit < 1kg T in the vacuum chamber)
Tritium losses in the fuel cycle
Wall material issues
…
Ref: K. Heinola et al., J. Nucl. Mater. 463 (2015) 961-965, doi:10.1016/j.jnucmat.2014.12.098
tile 0
tile 1
(Joint European Torus)
75% of retention in divertor

## Slide 5: JET retention
JET
Widdowson et al., PFMC 2021, Phys. Scr, 2021
Outer Divertor 16%
Inner Divertor 57%
Divertor 75%

## Slide 6: LID-QMS method
high energy laser (~20 J, 1 ms)
CFC substrate
Laser-Induced Desorption (LID)
Laser only heats the surface, no ablation, no crater formation, no plasma needed
Possible surface modifications due to heating :
roughening of surface
melting (only for high power)
minor evaporation
layer delamination due to thermal stresses in short time
Laser spot diameter ~3 mm
vacuum chamber
RGA (Residual Gas Analysis) e.g. by QMS (Quadrupole Mass Spectrometry)
H2, D2, HD, DT,…
Be melting point
Be layer thickness in JET is heated and thermally releases retained gases
substrate
Example of calculated temperature
W Mo
Be
Be
W
plasma deposit
Mo
CFC
Gas released by the laser heating is quantified → space resolved in situ diagnostic for fuel content

## Slide 7: Ex situ LID-QMS of JET divertor tile 0
Ref [1]: J. Likonen et al, Nuc. Mat. En. 19 (2019) 300-306 doi:10.1016/j.nme.2019.03.012
JET inner divertor tile 0 exposed in all plasmas of 2014-2016. Then ex situ LID-QMS performed in FZJ.
Good agreement of LID-QMS with TDS (Thermal Desorption Spectroscopy) except close to plasma heated side (due to D retention below laser heating depth) here: Be:D layers, previously also proven for
bulk W, a-C:D layers and a-B:D layers
constant D inventory
tile 0
tile 1
2c
4c
6c
7c
1c
3c
5c
LID-QMS: Ø3 mm laser spots
TDS of cored samples
5c low
3c low
1c low
LH14N

## Slide 8: In situ LID-QMS setup at JET
1
M1
M2
M3
M4
M5
Torus Hall:
Roof Lab above Torus Hall
Field Lens
movable mirror
cooledBeam Dump
Coaxial Camera
Fibre to pyrometers
Laser fibre
passive Beam Dump
Optics Table
Pyrometers
Total Transmission of Laser beam path: ~53%
Transmission of vacuum window: ~85% Light reflection on surface: ~25%
Total laser power for heating: ~34%
Diode-pumped Fibre Laser: Ytterbium 1071 nm, 25 kW (IPG YLS-25000)
Coaxial camera observation and temperature measurement of the laser spot
3 cylindric mirrors squeeze laser beam to be circular on the target
Ref: M. Zlobinski et al., Nucl. Fusion 64 (2024) 086031, doi:10.1088/1741-4326/ad52a5

## Slide 9: Laser accessible area
2.29 mm
3.06 mm
tile 1 ≈ 64° ≈ 13°
≈ 27°
tile 0 laser angle ≈ 48° (to surface normal)
Laser spot ellipticity show on front of tile 1:
with beam ellipticity settings for tile 0, PLaser=21.5 kW, 1 ms
Vacuum window
toroidal direction
poloidal direction
upper inner divertor of JET
142 mm

## Slide 10: Tritium Monitor Diagnostic at JET
Laser beam path length after laser fibre | 35.3 m | 27.8 m
Amount of Laser Mirrors | 12 + 1 beam splitter | 16 + 2 beam splitters
First mirror to target | 5.5 m | 5.8 m
Laser spot size on target |  3 mm |  5 mm
Laser power | 25 kW | 60 kW
Largest mirror dimension |  15 cm (M2, M3, M5) | 32 cm x 14 cm (M3)
In-vacuum/ex-vacuum | complete system ex-vessel | Mirrors M1-M6 in-vessel (in vacuum)
Scanning mirror to target | 5.5 m (M1) | 13.4 m (M7)
Neutron dose | ~ 1021 n (during last year of JET operation) | ~ 1027 n (for whole ITER lifetime)
Remote Handling compatiblity | No | Yes
Fast scanning mirror | 50 Hz stop-and-go, 3 mm spot-to-spot | max. 64 Hz, 3.5 mm spot-to-spot
Camera Observation of target in VIS | coaxial and through separate window | coaxial
Temperature Measurement | 3 single color point pyrometer | 3 single color point pyrometer, IR camera, spectrometer
Target illumination | Broadband VIS light (IVIS) + halogen lamp | Lasers of a separate system (GL1)
Special features | Ellipticity correction (M8, M9, M10), Focus length compensation (M2, M3) | Shutter to Plasma, Remote optical surface inspection by ultra-sort laser pulses, Remote laser position sensors

## Slide 11: Detection System
QMS5
Laser
QMS3=RGA3, QMS7=RGA7, …
QMS1=RGA10,QMS2=RGA11
sub-divertor
Octant 8
The gases desorbed from the laser spot are quantified by Residual Gas Analysis (RGA) mainly by Quadrupole Mass Spectrometers (QMS):
3 QMS in the main torus
4 QMS in the sub-divertor
JET torus
Detection of released gases by mass spectrometry in JET proven to be very sensitive
  if mass spectrometers are located close to the torus with good gas conductance (DT and T2 mainly detected by QMS 5, 1, 2)
  if background partial pressures are low (low wall out-gasing)  waiting time after plasma needed (> 1 night)
  if torus pumping is reduced such that mass spectrometer repetition time << pump-out time
Octant 8
LESSEONS LEARNED:

## Slide 12: Laser power scan to confirm complete D desorption
Before this measurement: poloidal and toroidal scan to identify area of const. D retention (green rectangle)
Laser power scan to determine minimum power for complete fuel desorption: 19 kW for 1 ms
19 kW
21 kW
Ref: M. Zlobinski et al., Nucl. Fusion 64 (2024) 086031, doi:10.1088/1741-4326/ad52a5

## Slide 13: QMS Calibration
Quadrupole Mass Spectrometer calibration by short (~1 s) D2 gas puff into the torus
Gas injection module volume V = 1.5 L
Pressure drop Δp in module after injection measured by baratron
N corresponds to the partial pressure jump in the mass spectrometer signal in the torus
Very low partial pressure jump is used, which is similar to the jump during the laser pulse (~1018 D atoms)
repeated on each measurement day to account for different wall loading, pumping conditions,…
injected gas molecules N
QMS5 calibration gas puff
QMS5 with laser pulse
Mass 4 (D2)
Ncalib
Time / s
Nsignal
Background of partial pressure on mass 4 subtracted from both

## Slide 14: LID-QMS evaluation by different methods
Integral method Integral of the gas puff
Jump method (Laura Laguardia) presentation on Thursday 9:00:
Integral method was rejected because data do not go back to initial value
Dynamic method works for the case of many data points in a smooth decay, but not for low signals with high noise. Jump method always works. So far the prime option for evaluation of LID-QMS measurements and calibration.
Dynamic method (Gabriele Gervasini) presentation on Thursday 9:25 “gas pumping method”:
∫
∫

## Slide 15: QMS Calibration
3 different D2 gas amounts puffed into torus, each amount repeated with slightly different amount → 6 gas puffs on the same day
Dp V / mbar L
QMS signal jump DI / A
Amount of puffed gas
QMS3
Linear fit
1s = 1.6%
Dp V / mbar L
QMS7
QMS signal jump DI / A
Amount of puffed gas
Factor of 5 lower sensitivity
Very good detector linearity in the range of the laser-induced signal
DI = K Dp V
Calibration factor K
Linear fit 1s = 1.7%

## Slide 16: Small improvement of Jump Method
1.) fitting of signal before and after jump
2.) extrapolate both functions to the time of laser pulse
3.) use the difference at the laser pulse as jump
function 1
function 2
laser pulse
difference

## Slide 17: Open Issue 1: QMS Cracking Pattern and overlapping masses
D2 He HT
DT
HD
T2
Deconvolution of hydrogen isotopologues e.g. HT and D2 both on mass 4. Complication only if T present. With H and D alone and a high-resolution QMS that can distinguish He from D2 there is no overlapping!
H2
H

## Slide 18: Open issue 1: QMS Cracking Pattern and overlapping masses
D2O
Open:
Deconvolution of cracking pattern inside QMS (mainly mass 1-6)
And overlapping masses (e.g. HT and D2)
Cracking of H2, D2 into atoms ~ few % only, thus negligible. But overlapping masses still an issue.
=> Existing deconvolution programme (by Drenik) to be understood
=> Take into account gas mixing due to species interaction time and calculate back with gas matrix (Sergienko, Matveev)

## Slide 19: Open Issue 2: LID-QMS calibration gas puffs: fractional injection
Calibration gas distribution in fractional volumes of the JET vessel (main torus, 2 x pumping tower) Gennady Sergienko (presentation on Thursday 10:30)
GIM4 used for calibration gas injection
Pumping Tower Oct 1
JET
main torus
Pumping Tower Oct 6
Laser in Oct 5
pumps

## Slide 20: Open Issue 3: calibration factors for mass 5 and 6
Deduction of calibration factor for mass 5 and 6 from mass 2 and 4
Mass Ks RGA3 (A/mbar*L) Ks RGA7 (A/mbar*L)
2 5.255419e-8 1.368379e-8
4 7.968536e-9 2.292932e-9
6 9.305809e-9 5.266764e-9
Ks RGA sensitivities #104589 H2, D2, T2 injections
Jump in current method Ks (A/mbar*L)
Measured QMS calibration factors
(G. Gervasini et al 2024 Nucl.Fusion64 126003)
Courtsey by D. Matveev
Expected from theory

## Slide 21: Open Issue 4: QMS Calibration Factor Variation
Open:
Do we need to systematically understand the different calibration factors on different days?
Or
Sufficient to use calibration factor of the same day as LID-QMS?
=> Maybe use first shot of each day (which was often in the same area) to find whether this is a variation only in calibration or also in LID-QMS.
Several reasons for variation identified:
RGA settings change (understood)
JET pumping change (could be compensated with dynamic method?)
Background pressure change (reason: non-linearity of QMS? But within each day gas puffs show good linearity)
Ref: Laguardia et al, PSI2024

## Slide 22: in situ Measurements with LID-QMS at JET inner upper divertor
14IWG1A
2019-23
14ING1C
2019-23
LH14W
2011-23
LH14N
2019-23
RH14W
2011-23
14ING1D
2011-23
RH14N
2019-23
14IWG1B
2011-23
Tile 0:
Tile 1:
Poloidal Scan with 1 laser pulse per position
9 laser pulses
25.9.2023 JPN 104423 -104441 cryo: LN2 only Turbo1
surface coordinate s = 0

## Slide 23: in situ Measurement of D by LID-QMS
Tile 0 (LH14N exposed 2019-2023), QMS 3
Assumptions: mass 4 is D2
mass 3 is HD
mass 5 is DT (i.e. neglecting
possible HT on mass 4 (but improbable because little HD and H2)
triple-ions like H2D+ on mass 4 or HD2+ on mass 5, which are improbable due to low partial pressure pD2 < 7E-8 mbar)
using the calibration factor of mass 4 also for mass 3 and mass 5
MEASUREMENT LIMIT for 20% uncertainty 2.6E18 D/cm2, 1.9E17 D atoms for a single spot
Detection limit on mass 4 (i.e. 100% uncertainty):
1.1E18 D/cm2, 8E16 D atoms for a single spot for this QMS, for this day (this background, pumping conditions, wall-outgassing etc.)

## Slide 24: Comparison to ex situ LID-QMS of JET divertor tile 0
JET inner divertor tile 0 (LH14N) exposed in all plasmas of 2019-2023. Then in situ LID-QMS performed in JET.
JET inner divertor tile 0 (LH14N) exposed in all plasmas of 2014-2016. Then ex situ LID-QMS performed in FZJ
Plausibility check:
qualitatively similar D profile (const. area + decrease towards plasma)
quantitatively:
4 years exposure with 41.9 h divertor plasma vs. 2 years exposure with 18.5 h divertor plasma (factor 2.4)
2.40.5 = 1.5__
RE Recently shown by Zayachuk et al* that total retention roughly scales with square-root of plasma exposure time
ex situ expected to be lower than in situ due to outgassing and exposure to atmosphere
plasma side
factor 1.7
* Zayachuk et al., NME 42 (2025) 101872, doi:10.1016/j.nme.2025.101872

## Slide 25: in situ Applications of LID-QMS in DT campaign
Monitoring of T retention growth due to plasma wall interaction
    4 x 3 cm2 rectangles “cleaned” from fuel by overlapping laser spots before tritium operation
    one rectangle measured every week in “Raster mode”: 99 laser spots every 3 mm every 20 ms to increase the desorbed area
    Additional weekly measurement of rectangle 1
4 cm
3 cm
4 cm
  4 x 3 cm2 “cleaned” rectangle overlapping with 1.5 mm spot-to-spot distance
  99 laser spots with 3 mm spot-to-spot distance
D2
DT
HD
T2
Deconvolution of overlapping masses: Work ongoing but in this case TH and HD2 have probably negligible contribution to mass 4 and 5 as there is little H2 (mass 2)
D2 , (TH) DT, (HD2) HD
Ref: M. Zlobinski et al., Nucl. Fusion 64 (2024) 086031, doi:10.1088/1741-4326/ad52a5
Measurement after 1 week of JET operation:

## Slide 26: in situ Applications of LID-QMS after DT campaign
Monitoring of T depletion by detritiation methods / Assessment of detritiation methods (Effectiveness? Where?)
    Monitor the tritium content during different clean-up techniques:
      Backing to 320° C
      ICWC plasmas
      Raised Inner Strike Point plasmas (RISP)
      Glow discharge plasmas
    Measurement by repetition of a poloidal scan of the inner divertor with LID-QMS shifted by few mm after each cleaning method
Example Measurement: LID-QMS on top of tile 1 before and after RISP plasmas
Laser pulse
Mass 5 (DT)

## Slide 27: in situ Measurements with LID-QMS at JET
Final Measurements at JET after last plasma in December 2023: Plan: direct comparison of ex situ to in situ LID-QMS in 70 positions ( ) on tiles retrieved from the torus and with LIBS
14IWG1A
2019-23
14ING1C
2019-23
LH14W
2011-23
LH14N
2019-23
RH14W
2011-23
14ING1D
2011-23
RH14N
2019-23
14IWG1B
2011-23
tile 0 (= HFGC)
tile 1

## Slide 28: LID-QMS comparison tiles with LIBS
14IWG1A
2019-23 Thursday
14ING1C
2019-23
LH14W
2011-23 Tuesday
LH14N
2019-23
RH14W
2011-23 Friday
14ING1D
2011-23
RH14N
2019-23
14IWG1B
2011/18? -2023 Wednesday
Tile 0: or HFGC
Tile 1:
Orange: Direct comparison of in-situ LIBS and in-situ LID-QMS possible
Green Tiles: retrieved from JET for comparison of in-situ to ex-situ LID-QMS
Compare with 4x4 LID-QMS:
Masses available: 2-6, 20, 28, 40 (RGA3) 2-6, 19, 20 (RGA5, RGA11) 1-6, 19, 20, 28 (RGA7)

## Slide 29: in situ Measurements with LID-QMS at JET
Laser Shot Plan (old photo overlaid with virtual laser spots):
14IWG1A
2019-23
14ING1C
2019-23
LH14W
2011-23
LH14N
2019-23
RH14W
2011-23
14ING1D
2011-23
RH14N
2019-23
14IWG1B
2011-23
Hi-Resolution photography on Mascot robotic arm:
4 Tiles 0 retrieved and being disassembled
tile 0 (= HFGC)
tile 0 (= HFGC)
tile 1

## Slide 30: Backup-Slides

## Slide 31: Poloidal Scan in situ on tile 0 (LH14N)
14ING1D
2011-23
Absolute Calibration:
Calibration factor Km4 from D2 puffs on the day of experiment
Take the same factor for mass 3 and 5 ? Or interpolate result of T2, D2, H2 puff ?
Chart: tile 0 (LH14N)
| 70.75 | 100.75 | 110.75 | 120.75 | 130.75 | 140.75 | 150.75 | 80.75 | 90.75
mass 4 | 1.7915e-07 | 2.0065e-07 | 2.0065e-07 | 2.2931e-07 | 2.2216e-07 | 1.8632e-07 | 1.4332e-07 | 2.7948e-07 | 2.0781e-07
mass 3 | 1.0033e-07 | 1.2183e-07 | 1.2899e-07 | 9.316e-08 | 1.2181e-07 | 1.0748e-07 | 9.315e-08 | 1.2899e-07 | 1.2183e-07
mass 5 | 3.583e-08 | 6.45e-08 | 2.866e-08 | 5.017e-08 | 5.015e-08 | 3.583e-08 | 2.149e-08 | 5.015e-08 | 6.449e-08
Assumption: mass 3 is HD, mass 4 is D2, mass 5 is DT (i.e. neglecting possible HT on mass 4 and triple-molecules like H2D on mass 4 or HD2 on mass 5, which are improbable due to low partial pressure)
Total D = DI3 * Km3 + 2 * Km4 * DI4 + Km5 * DI5
RGA3 jump / A
Ks RGA3 2.0185e-5 A/mbar*L (± 7.142e-9)
DI3, DI4, DI5
Km3
Km4
Km5

## Slide 32: Extrapolation to JET and ITER
Fusion facility | Volume V [m3] | Pumping speed S [l/s] | Pump time tpump [s] | Back-ground pD2 [mbar] | Inner surface A [m2] | D2 outgasing rate D atoms /s | Lower detection limit D-atoms 1016
TEXTOR | 17.2 | 2100 | 8.2 | 1.6 10-7 | 38 | 1.6 ×1016 | 1.5
JET | 200 | 8000 (turbo) | 25 | 3 10-7 | 200 | 1.17× 1017 | 20
ITER | 1400 | 75000 | 18 | 1.3 10-7 | 1000 | 5 x1017 | 130
Strong dependence of the lower detection limit on the volume
Ref: A Huber et al 2001 Phys. Scr. 2001 102

## Slide 33: RGA 3 raw data