## Slide 1: Introduction to LIBS on JET, including type of data available, progress on results
Jari Likonen

## Slide 2: LIBS in fusion reactors
Laser-induced breakdown spectroscopy (LIBS) is a type of atomic emission spectroscopy which uses a highly energetic laser pulse as the excitation source
Laser is focused to form a plasma, which atomizes and excites species
These excited-state species emit light at unique wavelengths
This light can then be collected with a spectrometer
LIBS allows detailed inspection of vacuum vessel of fusion reactor
  Particularly critical is to assess the amount of radioactive tritium retained in the layers
  Also possible to determine thickness and elemental composition of the co-deposits
  LIBS provides an (almost) unlimited access to different regions of reactor wall – without the need to remove wall components from the reactor
  LIBS parameters can be tailored such that impact on wall structures is minimized 
  no need for replacement after LIBS measurements

## Slide 3: LIBS system at JET
LIBS enclosure developed by ENEA and optimised by FZJ
LIBS enclosure mounted on the JET MASCOT arm
RH boom moves LIBS tool close to the wall, outermost cone is in contact with tile to be investigated
Argon gas flux applied to increase signal to noise ratio
Optical fibre guides emitted light into spectrometers, located outside of the vessel

## Slide 4: In-vessel LIBS applications at JET
Optics Table
2.Littrow spectrometer: F#~4,resolution~30 pm(slit 100um), wavelength range 4 nm- Hα detection
1
2
3
4
4. PMT :monochromator +3 different filters- Hα detection
1. Echelle : F#10,resolution~70pm, 200-750 nm- T calculation
3. Compact C-T spectrometer: resolution 0.5 nm, 300-1100 nm- Layer information
Fiber bundle
Laser box
4 spectrometers
JET-LIBS setup
1
2
3
4

## Slide 5
X
Z
X
Z
X
Z
X
Z
X
Z
X
Z
X
Z
X
Z
4D
VW
Oct5
VW
Oct1
LIBS – Laser Induced Breakdown Spectroscopy – TOROIDAL MAP
W – coated CFC
Bulk W
Bulk Be
Be coated inconel
Inconel (vacuum vessel)

## Slide 6: LIBS full distribution map
W – coated CFC
Bulk W
Bulk Be
Be coated inconel
Inconel (vacuum vessel)
Poloidal direction
Toroidal direction
402

## Slide 7: LIBS at JET in 2025
1st Data analysis meeting (Feb 6th, FZJ)
Reported in March, PWIE meeting in Prague
PFMC
Several contributions (Likonen, Almaviva, Gasior, Yi)
Poloidal distributions of H isotopes, Be, W
11th Data analysis meeting (Oct 31st, Teams)
JET LIBS calibration pulses (pure materials)
  Be calibration: 2XR8 pulse #492 (from C3 tile)
  Inconel (vessel wall) OCT1 middle: pulse no. #357
  Pure W (W-wedge – Lamellae): Stack A pulse #20 and repeated pulse (with setting changes) pulse #765
  W-CFC Top of tile 8: pulse no. #51 (250 shots)
13th Data analysis meeting (Nov 28th, Teams)
Discussion on the joint LID/QMS-LIBS meeting

## Slide 8: FZJ: Littrow results from Module 14(LID-QMS positions)
LIBS pulses: #xy
LID-QMS pulses: #105abc
LID-QMS and LIBS analysis locations on tiles 0 and 1

## Slide 9: FZJ: Littrow results from Module 14 (LID-QMS positions)
2D map for (D+T)/(H+D+T)
surface
HFGC
Tile 1

## Slide 10: FZJ: Littrow results from Module 14 (LID-QMS positions)
2D map for (D+T)/(H+D+T)
surface
D+T amount not constant toroidally
HFGC LH14W
HFGC RH14W

## Slide 11: Aryelle: Be profile on divertor
1
LBSRP_14W LH
C1 & C21
14BWG4A
14ON G7A
6
7
0
3
4
5
8
1
HFGC_LH 14W
14IWG1B
14IWG3A
14ON G8A
Heaviest Be deposition on HFGC tile
and top part of tile 1
See Almaviva O1
Divertor module 14 with LIBS locations

## Slide 12: Aryelle: D profile on divertor
0
1
3
4
5
7
8
Divertor module 14 with LIBS locations
1
LBSRP_14W LH
C1 & C21
14BWG4A
14ON G7A
6
7
0
3
4
5
8
1
HFGC_LH 14W
14IWG1B
14IWG3A
14ON G8A
H isotopes mainly retained on HFGC tile
and top part of tile 1
Co-deposition together with Be
See Almaviva O1
T-D-H slowly descreasing

## Slide 13: Littrow: D and T map for divertor
Overall T amount low along divertor tiles
See Yi 4889 POB-10, Thursday
1
20
12
15
16
17
8
9
10
11
4
3
0
1
3
4
5
7
8
0
1
3
4
5
7
8

## Slide 14: Calibration free LIBS
CF procedure (A. Ciucci, Appl. Spec. 53(8), (1999), 960-964) aims to quantitatively estimate chemical elements detected in LIBS plasma.
CF will be applied to estimate the concentration of T+D+H with respect to W and Be.
CF is based on experimental intensities of the emission lines.
If LIBS plasma can be considered in local thermodynamic equilibrium (LTE) these intensities can be expressed as follows:
where, 𝐼ki𝜆 = exp. intensity of the k i transition, N𝑠 = concentration of species, 𝐴𝑘𝑖 = transition probability for given line, gk is k level degeneracy, 𝐸𝑘 the upper energy level of transition, 𝑘𝐵 Boltzmann constant, T is the plasma temperature, Us(T ) is the partition function for emitting species at plasma temperature Te.
Taking the logarithm of equation above we get (Boltzmann plot)
where

## Slide 15: Calibration free LIBS
Slope m of plots is related to plasma temperature, while the intercept qs is proportional to logarithm of the species concentration.
Boltzmann plot method provides the temperature from only one ionization stage.
More general analysis is obtained using Saha-Boltzmann equation which provides relationships between transition line intensities from different ionization stages.
qs(Al I)
qs(Mg II)
qs(Mn II)

## Slide 16: ENEA: Initial results of spectral analysis at JET, composition of co-deposited layers and fuel content: example: point 357 (Inconel reference). In-depth composition
500
LIBS spectrum of Inconel
Depth profiling of Be, W, Cr, Ni, T-D-H

## Slide 17: ENEA: Initial results of spectral analysis at JET, composition of co-deposited layers and fuel content: example: point 357 (Inconel reference). Fuel content
500
Intensity ratio of the Balmer alpha band
(T+D+H) with respect to Cr I line at
425.44 nm and Ni line at 310.16 nm
Relative atomic concentration of Be, W, H, Ni,
C by applying the closure relation (sum of con-
centration = 100%)

## Slide 18: ML models development at the IPPLM – overview on results
Task and questions to be addressed
IPPLM specific: Implementation of machine learning and digital signal processing algorithms to facilitate data analysis
Approach
Application of digital data filtering for preliminary data engineering
Application of DNN-based autoenconer for dimensionality reduction and denoising
(DNN=Deep Neural Networks)
Application of DNN and CNN for data analysis (CNN=Convolutional Neural Networks)
Results
Autoencoder
Input spectrum
Output spectrum
Profile reconstruction
14ONG8A
Line reconstruction, Be I 457.3nm

## Slide 19: ISSP UL: Results of spectral analysis, composition of co-deposited layers and fuel content
Task and questions to be addressed
Spectral analysis of Aryelle data
Approach
CF-LIBS has been applied to the selected JET tiles (#492 Be ref. 2XR8; #20 W ref. 14WLH
AR3; #357 Inconel ref. Oct1 Middle)
Results
Transmission multiplier and echellogram correction have been applied to adjust
intensities of spectral lines
For calculation of electron temperatures, inverse line intensity (wi=1/Ii, wi=1/Ii2)
has been used as a weighting factor to reduce influence of self-absorbed spectral lines.
Weighted Boltzmann plots, in general, yielded slightly lower electron temperatures
Accuracy of relative atomic concentrations strongly depends on determination
of electron temperature.
Electron densities yielded 6…150 x 1021 m-3
Depth profiles and Te for the selected JET tiles
Red: unweighted Te , green: weighted Te
W ref. 20 14WLH AR3 C1
Be ref. 492 2X08 R7 C3 C7
D/H and Be signals appear up to the 40th pulse, exhibiting slow decay. He contamination to the surface is small
No significant amount of impurities was detected on the surface

## Slide 20: Relative concentrations on species in studied samples
Sample | Emitting species | Electron temperature, K | Relative concentration (averaged over all relevant lines), %
W (#20) | W I | 9573 (W I) | 99.0 ± 1.0
 | H I |  | 1.0 ± 1.0
Be (#492) | Be I | 11807 (Be I) | 98.5 ± 1.0
 | H I |  | 1.5 ± 1.0
Inconel (#357) | Ni I | 10267 (Ni I) | 56.27 ± 0.85
 | Cr I |  | 40.40 ± 0.65
 | Fe I |  | 3.32 ± 0.68
 | H I |  | (-1.43 ± 0.35)∙10-5
Notes: Ielikt grafiku, kur nekas nemainās

## Slide 21: UT: Analysis of data from LIBS JET campaign - reference samples
#357 Inconel OCT1 Middle
Chart: LIBS shot nr.
Chart: Energy (eV)
| 4.1539999999999999 | 4.2300000000000004 | 4.5380000000000003 | 4.1050000000000004 | 4.0880000000000001 | 3.6549999999999998 | 3.64 | 3.706 | 3.7959999999999998 | 3.06 | 3.66 | 3.6349999999999998 | 3.5419999999999998 | 3.6349999999999998 | 6.01
 | -19.7265 | -19.7119 | -20.1854 | -19.9534 | -19.5578 | -18.5394 | -18.7347 | -18.9803 | -18.4158 | -18.2839 | -18.1676 | -18.8305 | -22.4864
 | -20.667 | -21.0291 | -20.7427 | -21.5097 | -22.3448 | -20.9395 | -18.5492 | -19.4185 | -19.2535 | -18.7453 | -19.0482 | -19.3969 | -17.6726
 | -20.0629 | -19.8013 | -19.6281 | -20.5161 | -20.5534 | -20.6317
 | -14.9291 | -17.8387 | -16.6045
 | -20.0921
Chart: Energy (eV)
| 4.1539999999999999 | 4.2300000000000004 | 4.5380000000000003 | 4.1050000000000004 | 4.0880000000000001 | 3.6549999999999998 | 3.64 | 3.706 | 3.7959999999999998 | 3.06 | 3.66 | 3.6349999999999998 | 3.5419999999999998 | 3.6349999999999998 | 6.01
 | -18.2355 | -17.9678 | -18.5452 | -17.6098 | -18.22 | -16.7312 | -16.9923 | -17.6151 | -16.915 | -17.1105 | -17.0146 | -16.703 | -19.6347
 | -18.4982 | -18.1073 | -17.8555 | -18.1259 | -18.3816 | -18.4852
 | -17.2827 | -20.0058 | -20.1257
 | -21.7016
 | -15.9734 | -17.1784 | -18.3834 | -19.5884 | -20.7934
 | -15.713 | -16.918 | -18.123 | -19.328 | -20.533
 | -14.3811 | -18.0051 | -21.6291
Shots 1-10: temperatures from W I, Ni I and Cr I lines are comparable
Shots 291-300: CF-LIBS Cr/Ni ratio 0.26 in the bulk sample is consistent with Inconel
Te = 0.57-0.63 eV
Te = 0.83 eV
Boltzmann plot

## Slide 22: UT: Analysis of data from LIBS JET campaign - reference samples
#20 Pure W LBSRP 14WLH
Chart: LIBS shot nr.
Chart: Energy (eV)
| 3.46 | 3.43 | 3.41 | 3.79 | 5.36 | 5.04 | 3.41 | 5.24 | 5.36 | 4.84 | 3.69 | 3.27 | 3.25 | 3.25 | 4.74 | 2.66 | 3.2469999999999999 | 2.97 | 2.97 | 3.31 | 4.3499999999999996 | 3.25 | 3.07 | 2.97 | 2.66 | 13.119 | 13.414 | 13.364000000000001 | 12.344000000000001
 | -10.1502 | -9.1741 | -9.729 | -10.1175 | -11.4053 | -11.4593 | -7.5751 | -11.7656 | -10.9206 | -11.2474 | -9.5129 | -8.5776 | -9.58 | -9.501 | -11.3758 | -7.8485 | -9.2648 | -7.311 | -8.7069 | -8.9472 | -9.8574 | -8.7077 | -8.6855 | -8.5402 | -7.7519 | -21.1353 | -22.754 | -22.2353 | -20.5001
 | -14.488
Saha-Boltzmann plot (W,D+H)
Chart: LIBS shot nr.
| 1-15. | 16-30. | 31-45. | 45-60.
Saha-Boltzmann | 0.76511 | 0.76681 | 0.76394 | 0.77024
Boltzmann | 0.7687 | 0.7458 | 0.7483 | 0.7963
Chart: LIBS shot nr.
| 1-15. | 16-30. | 31-45. | 45-60.
 | 33500000000000000 | 26200000000000000 | 29200000000000000 | 32100000000000000
Depth profiles are quite noisy and averaging of spectra over 10-15 LIBS shots gives more reliable results for CF-LIBS
Temperature from Boltzmann plot of W I lines and Saha-Boltzmann plot of W I and W II lines consistent
Electron density determined from Ar I and Be I lines consistent
H isotope quantification requires additional work

## Slide 23: LIBS data access – STEP 4
On your PC/desktop create a shortcut folder:
Username: local\nobody
Password: <blank>(don't put any, leave empty)
\\dstorehost.JETdata\rh-libs
For next prompt use:

## Slide 24: Concluding remarks
LIBS experiment at JET was successfully completed in October 2024 (840 locations)
LIBS data analysis progressing
  Data analysis meetings on a monthly basis
  Good progress in CF LIBS
Several contributions in PFMC
Additional reference LIBS measurements including some DTE3 samples to be made at VTT (~March-April 2026) depending on availability of Aryelle, Littrow and DEMON (LTB Lasertechnik Berlin GmbH) spectrometers.
Capabilities of DEMON spectrometer (not used at JET) will be tested (separation of D-T components) at VTT.