## Slide 1: Tile HFGC LH14W (2011-2023) LIBS retention (T, D, H, other)
Salvatore Almaviva
ENEA, Italian National Agency for New Technologies, Energy and Sustainable Economic Development
Joint LID-QMS/LIBS meeting, ENEA-Frascati, 9-12 December 2025

## Slide 2: Summary
Wavelength and intensity calibration
Calibration free LIBS (CF) procedure JET reference points
CF-LIBS on JET tile HFGC LH14W (2011-2023)
Conclusions and next steps

## Slide 3: LIBS analysis of points 59-60-61-62 (HFGC LH14W)
The analysis area is located in octant 5 of JET and is composed of W - coated CFC PFCs
W – coated CFC
Bulk W
Bulk Be
Be coated inconel
Inconel (vacuum vessel)
Courtesy by UKAEA (I. Iepu)

## Slide 4: LIBS analysis of points 59-60-62-63 (HFGC LH14W)
HFGC LH14W (2011-23) |  | 
LID-QMS #no | LIBS #no | Comment
105852 | 63 | Poloidal assessment
105853 |  | 
105854 | 62 | 
105855 |  | 
105856 | 60 | 
105858 | 61 | Toroidal assessment
105859 |  | 
105860 |  | 
105855 | 60 | 
105866 |  | 
105865 |  | 
105864 |  | 
105867 | 59 | 
HFGC
LH14W (2011-23)
852
853
854
855
856
858
859
860
861
864
865
866
867
63
62
61
60
59
Toroidal
Poloidal

## Slide 5: Wavelength calibration
The software package developed by UT was used to perform the wavelength calibration of the Echellograms acquired during the experimental campaign at JET.
The position of some emission lines has been checked after the calibration: residual shifts are in the range of few tenths of pm
Element | Wavelength (nm) | Element | Wavelength (nm)
Be II | 313.042 | Be I | 457.27
Be II | 313.1 | Ar I | 696.54
Be I | 332.1 | Ar I | 706.72
Be I | 332.13 | Ar I | 727.3
W I | 400.875 | Ar I | 738.4
W I | 407.44 | Ar I | 742.23

## Slide 6: Intensity calibration
The optical transmission file was implemented to rescale the intensity of the LIBS spectra, by taking into account all the optical components of the LIBS system. The spectral region 344-385 nm was ignored because of the low signal related to the low optical transmission of the whole system
×
=

## Slide 7: CF-LIBS procedure
The CF-LIBS algorithm was developed in a MATLAB environment following the next steps:
1) Automatic identification of the peaks in the experimental spectrum using the "findpeaks" function
2) Automatic assignment of the newly found peaks to the atomic and ionic emission lines of the chemical elements in the PFCs of JET (e.g., Be, W, Mo, Ni, Cr, Fe, C) based on the NIST database, comparing the relative position of the peaks.
3) Calculation of the electron density of the LIBS plasma from the line broadening of some reference lines of the above-mentioned chemical elements using the "peakfit" function
4) Calculation of the electron temperature of the LIBS plasma using the extended Boltzmann plot of the above-mentioned chemical elements
5) Calculation of the H-D-T band related to the Balmer alpha emission of the three H-isotopes
6) Application of the closure relation and calculation of the element concentrations in atomic %

## Slide 8: CF-LIBS procedure: identification of the peaks in the spectrum
The “findpeaks” function returns a vector with the local maxima (peaks) of the input signal. The parameters needed to implement the function are:
1) minimum height of the peak (set as a fraction of the maximum signal).
2) prominence (reducing any possible interfering peak)
3) minimum spectral distance between the peaks (in nm).
4) peak width (in nm)

## Slide 9: CF-LIBS procedure: identification of the peaks in the spectrum
The peaks identified in the spectrum are identified as atomic and ionic emission lines of the chemical elements of the PFCs of JET (e.g., Be, W, Mo, Ni, Cr, Fe, C) comparing their wavelength with those of the NIST database.
The spectroscopic parameters to apply the calibration free procedure (Energy of the upper level, Einstein’s coefficient, degeneracy of the transition) are the obtained by the NIST database.
Exp wv (nm)
NIST wv (nm)
Difference (nm)
Peak area
NIST gkAki
ln(Il/gkAki)
Ek (+ Eion)
Atom/Ion
W Ion
W atom
C Ion
Ni atom
Be ion

## Slide 10: CF-LIBS procedure: calculation of the electron density
3) The plasma electron density is calculated by using the line broadening of some reference lines of the above-mentioned chemical elements using the “peakfit” function.
The peakfit function compute the Lorentz width of some reference lines useful for the electron density calculation.
Currently four lines are used in the peakfit function:
W I @ 426.94nm, W I @ 429.46nm, W I @ 430.21nm, Be I @ 457.27nm

## Slide 11: CF-LIBS procedure: calculation of the electron temperature
4) The electron temperature of the LIBS plasma is calculated by using the extended Boltzmann plot of the above-mentioned chemical elements: given the large number of atomic and ionic transitions the W I and W II atomic and ionic lines are generally used for the calculation of the electron temperature.

## Slide 12: Determination of the signal of the H-D-T band (Hydrogen isotopes)
Calculation of the H-D-T band related to the Balmer alpha emission of the three H-isotopes: the intensity of the H-D-T band (considered as a unique emission band) is obtained by the experimental H-D-T integral signal between 655.9 – 656.4 nm subtracted by the nearby integral noise signal between 655.4 – 655.9 nm (no emission lines have been detected in that spectral region in the LIBS spectra). If the H-D-T signal is lower than the noise the H-D-T signal is assumed to be zero (no Hydrogen isotopes)

## Slide 13: CF-LIBS procedure: test on reference points
Application of the closure relation and calculation of the element concentrations in atomic %: the CF-LIBS procedure and the closure relation is applied for each series of spectra.
492_2X8_R7C3C7_2X08 - R7 C3-C7 Be reference
357_VW1_M_Oct1_MIDDLE Inconel reference

## Slide 14: CF-LIBS procedure: test on reference points
6) Application of the closure relation and calculation of the element concentrations in atomic %: the CF-LIBS procedure and the closure relation is applied for each series of spectra. Results for the reference points (Te 15500 K, ne = 3*1017)
360 - 02ONG8B - R4 C2 W_CFC_reference
20_14WLH_AR3C1 PURE W_reference

## Slide 15: CF-LIBS procedure: LIBS points at JET 59-60-61-62-63
Application of the closure relation and calculation of the element concentrations in atomic %: points 59-60-61-62-63
59_LH14W_R1C1
60_LH14W_R1C2
59
60

## Slide 16: CF-LIBS procedure: LIBS points at JET 59-60-61-62-63
Application of the closure relation and calculation of the element concentrations in atomic %: points 59-60-61-62-63
59 R1C1
60 R1C2
61 R1C3
62 R1C4
63 R3C2
61
60
59
62
63

## Slide 17: CF-LIBS procedure: LIBS points at JET 59-60-61-62-63
Adjacent average over 5 points for W and T-D-H
Atomic conc. (%)
Atomic conc. (%)
Atomic conc. (%)

## Slide 18: Conclusions and next steps
1) An automatize sw procedure is being developed to process the large number of spectra acquired in JET and apply the CF-LIBS in short times, giving the concentrations of the detected species in atomic %
2) First results for the Hydrogen isotopes, considered as a whole, suggest a residual atomic concentration of 5-10 % in the surface layer
3) Be is detected at very high concentration in the surface deposits
4) The thickness of the surface deposits depends on the position of the analysis point
5) Next steps:
  update the procedure to include other elements (Ni, Cr, Mo…)
  Optimize the number of emission lines included in the procedure retrieved from the NIST database