## Slide 1: Commissioning of ps-LIBS on BiGyM and first tests with boron-coated samples
L. Laguardia, A. Cremona, S. Almaviva, S. Cipelli, A. Uccello, F. Cani
WP PWIE SP F kick-off meeting for 2026 activities

## Slide 2: Deliverable 2026
Deliverable ID | Deliverable Title
D007 | Complete installation and optical integration of LIBS on BiGyM and perform system calibration with test samples (ENEA)
Deliverable Owner | Beneficiary | PM | Deliverable (Team)
L. Laguardia | ENEA | 4 | L. Laguardia, S. Almaviva, A. Cremona
Human Resources 2026
Device | Beneficiary | Days | Related Deliverable
GyM/BiGyM | ENEA | 5 | D007
Machine Resources 2026

## Slide 3: Status of activities
ps-LIBS has been installed and tested in Diagnostic Laboratory at ISTP RU
Nominal λ [nm] | 1064
Pulse duration [ps] | 29 ± 5
Max pulse energy [mJ] | 70
Repetition rate [Hz] | 10
Spectral resolution (nm) | Up to 0.064
Focal lenght [mm] | 320
Min slit width [μm] | 10
Compact design
Correction of optical aberrations
Set of 3 interchangeable gratings
PMAX Intensified CCD camera (Gen II SR intensifier)
S. Cipelli curtesy
L. Laguardia | WP PWIE SP F kick-off meeting | 23 June 2026
Hα, Dα

## Slide 4: Status of activities
BiGyM linear device laboratory
Diagnostics laboratory
Hole for the optical path
Hole for the optical path
EKSPLA laser
BiGyM pre-chamber
S. Cipelli curtesy
Over the past year, alongside the GyM→BiGyM upgrade, the Diagnostics Laboratory was reorganized and equipped with all required safety systems to host LIBS and enable laser beam delivery from the Diagnostics Lab to BiGyM.

## Slide 5: ISTP-ENEA PLAN 2026 Commissioning of an in-situ LIBS on the BiGyM linear plasma device
Technical drawing of an optical path connecting the laser system to the BiGyM pre-chamber
Side view
~8 m optical path length
Beam enclosed by safety tubes
Use of 3 fixed mirrors + 1 motorized steering mirror for beam position scanning
Focusing lens
Lens + optical fiber for collection of plasma light
S. Cipelli and F. Cani curtesy
Diagnostic
Laboratory
BiGyM laboratory
ps-laser
source
BiGyM sample
Exhange chamber

## Slide 6
ps Nd:YAG
Princeton spectrometer
Sample Holder
sample
Motorized Mirror
Focusing optics
Plasma plume
ICCD
z-θ-φ translation stage
Optical fiber
Laser beam guide
BiGyM Top view
Tantalum shield
z
φ
θ
Collection optics
x-y scanning
X mm
1
10
3
4
5
6
7
8
9
2
Y mm
1
2
4
3
5
6
8
9
7
10
ps-LIBS optical layout and X–Y scanning

## Slide 7: On-going activities
Procurement of:
safety tubes
mirrors
motorized steering mirror
lenses
fiber‑bundle
Time-line
Installation
July
August
September
October
November
December
Procurement
Summer holydays
Testing with H. Wu and S. Almaviva
Testing
Results

## Slide 8: Boron-coated samples
Where will the boron samples we will use for the first tests come from? From FZJ, from IAP, or from somewhere else?
For the first tests with H. Wu and S. Almaviva, we would prefer the samples to already contain D, as it remains to be confirmed whether BiGyM will be ready for PWI experiments in October (although it is scheduled to become available by the end of 2026), leaving open the question of whether deuterium loading onto boron samples will be feasible in time.
5 Flush-mounted "PSI-2" samples
Tantalum
Rear-side
thermocouples